## Slide 1: Take Two Software Updates and See Me in the Morning: The Case for Software Security Evaluations of Medical Devices
Steve Hanna1, Rolf Rolles4, Andres Molina-Markham2, Pongsin Poosankam1,3, Kevin Fu2, Dawn Song1
University of California – Berkeley1, University of Massachusetts Amherst2,
Carnegie Mellon University3, Unaffiliated4

## Slide 2: Changing Medical Device Landscape
Increased software complexity
Software plays an increasing role in device failure
  2005-2009 (18%) due to software failure, compared to (6%) in 1980s
Increased attack opportunities
Medical device hardware and software is usually a monoculture within device model
Health Data
Connected
Devices
Medical
Device
28,000 adverse event reports in 14 Models recalled 2005-2010.
Automated External Defibrillators

## Slide 3: To be clear…
AEDs
ICDs
X

## Slide 4
Wisconsin requires daycare providers to be AED proficient
Global Automated External Defibrillators (AED) Market: Demand to Drive Growth; June 2009 U.S., European and Japanese External Defibrillation (PAD) Market Report. Frost & Sullivan. 2000. Valenzuela TD, et al. N Engl J Med. 2000;343:1206-1209. Caffrey S, et al. N Engl J Med. 2002;347:1242-1247.
The Population of AEDs Has
Increased Significantly Over the Past 5 Years
Automated External Defibrillator Milestones
AEDs Worldwide
1,582,691
1996
1998
2000
2002
2004
2006
2008
First AED with biphasic waveform
First save on US airline
74% survival rate in casinos
75% survival rate in O’Hare Airport
PAD Trial Published
New York requires AEDs in public places

## Slide 5: Our Objectives
Explore state of AED software security
Examine for standard software security flaws
  Data handling, coding practices, developer assumptions
Give insight into state of medical device software and potential for future abuse

## Slide 6: Desirable Medical Device Properties
The device should:
  Ensure that software running on a system is the image that was verified
  Detect compromise
  Verify and authenticate device telemetry
  Be robust: defenses and updates weighed with risks to patient

## Slide 7: Case Study
Analyzed Cardiac Science G3 Plus model 9390A
Performed static reverse engineering using IDA Pro
  Analyzed: MDLink, AEDUpdate and device firmware
Analysis using BitBlaze architecture
  BitFuzz, the dynamic symbolic path exploration tool
Remarks
  Problems likely not isolated to the G3 Plus
  Potential for abuse as devices become more connected

## Slide 8: Vulnerabilities Discovered
AED Firmware - Replacement
AEDUpdate - Buffer overflow
AEDUpdate - Plain text user credentials
MDLink - Weak password scheme
Vulnerabilities were verified on Windows XP SP2.

## Slide 9: Firmware Replacement
Firmware update uses custom CRC to verify firmware
Modified firmware, with proper CRC, is accepted by AED and update software
Impact: Arbitrary firmware
DEVICE COMPROMISED

## Slide 10: AEDUpdate Buffer Overflow
During update device handshake, device version number exchanged
AEDUpdate improperly assumes valid input
Enables arbitrary code execution
  Data sent from AED can be executed as code on the host PC

## Slide 12: Improving Medical Device Security for Developers
Lessons and open problems from the CS G3 Plus
  Cryptographically secure device updates
    No security through obscurity, ensures firmware authenticity
  Device telemetry verified for integrity and authenticity
    Defensively assume that data is not trusted
  Passwords cryptographically secure and easily managed
    Private data and life critical functionality should be protected by well-established cryptographic algorithms
  Defenses and updates weighed with risks to patient
    Medical devices should fail open

## Slide 13: Recommendations
Ensure the update machine is secure
  Physical isolation, virtual machine for fresh install
Follow FDA guidelines and advisories
Remain vigilant
  Monitoring physical access, routinely updating afflicted devices, and monitoring advisories released about the device

## Slide 14: Final Recommendation
We recommend continued use of AEDs because of their potential to perform lifesaving functions.
The attack potential is currently unmeasured and currently, these devices overwhelmingly save more lives than they imperil.

## Slide 15: Thank You
Questions?
  Contact:
    Steve Hanna (sch@eecs.berkeley.edu)
    Dawn Song (dawnsong@cs.berkeley.edu)
    Kevin Fu (kevinfu@cs.umass.edu)
  secure-medicine.org